Lab report   06/15/10   by QQ

General:

Chris is off for half week.

Equipment:

- Pump was ordered, will come at the end of this week(QQ)

- 100 ohm Potentiometer has been ordered, and received. They replaced the original 25 ohm potentiometer. So that two power supplies could provide heat for two chips at the same time.(QQ)

FLU Samples:

Run the upthawed sample from BUMC in sequence.

Can we run thawed samples?

Chip fabrication:

- QQ has checked AJ’s bonding.

- Thermocouples are expensive. QQ help pretreat them to get a good shape and then give AJ to bond.

- AJ make 4 chips a day. Half of them not well bonded. (success ratio from 1/3 to 1/2)

- Fabrication speed is slow. We are not getting 4 good chip each day for experiment.

Solution: a. get the priority to use the Carver. Notify out users come on weekend???(if possible)

b. chris help fabrication.

Experiment:

- Nanoport get off during SPE. (5 chip had the same issue one day)

Solution:

use a fixture.

Old SPE fixture is used now for the SPE part only. New fixture for the integrated chip will be made during the next several weeks.

- We run 6 samples last week. (SPE by Branden, PCR 1 by chris,  3 by QQ, 2 by Branden and QQ)

3 positive, 1 negative, 2 have not detected by bioanalyzer.

- QQ trained Branden to run two chips at the same time.

QQ will train chris to run two chips when he come back.