Fluidigm Genotyping Protocol Using SNPtype Assay on the 96.96 IFC

Preparing SNPtype Pool for Pre-AMP

*Note, make a 600 uL mix if testing out assay panels.

*Multiply total volume by number of plates if making larger working stock

- Use multichannel to pipette assays into a clean trough, add Low TE, and transfer to a 1.7 ul tube and vortex.

Specific Target Amplification Primers (STA Assay stock plate of 100 uM)	2.4 ul each (x 96)

Locus-specific Primers (LSP Assay stock plate of 100 uM)			2.4 ul each (x 96)

Low TE					         					139.2 ul_____________

600 ul

Preparing SNPtype Assay Stock Plates

- Use the multichannel to pipette the ASP Assays and LSP Assays into a stock plate, add Low TE (* Assays are well specific!)

Allele-specific Primers 1&2 (100uM ASP1 & 100uM ASP2)		3 ul each **well specific

Locus-Specific Primers (100uM LSP)					8 ul each **well specific

Low TE					         				29 ul per well ______________										40 ul (Split into 2 aliquots)

SNP Pre-Amp

*After feather extraction using KingFisher 96, you should end up with 93 samples + 1 negative control. Assign NTC’s to A1 & B7

- Add 5 ul of DNA to each well of the “PreAmp Plate”, dry down samples completely then add Qiagen MM + SNPType Pool Primers (LSP+STA). *This can take up to 1 hour

Pre-Amp Master Mix:

|  | Reactions |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Number of Reactions | 1X | 110X |  | 220X |  |
| Qiagen 2X PCR MM (uL) | 5 | 550 | uL | 1100 | uL |
| Primer Pool [0.2uM each/ reaction] (uL) | 5 | 550 | uL | 1100 | uL |
| Total (uL) | 10 | 1100 | uL | 2200 | uL |

•	Add 10 ul/well of PreAmp MM to dried down DNA, seal with PCR film and spin down

| Thermocyler: SORK FLUIDI PreAmpM |  |  |  |
| --- | --- | --- | --- |
| Cycle Step | Temp | Time | Cycles |
| Initial Denaturation | 95°C | 15 minutes | 1 |
| Denaturation | 94°C | 30 seconds | 14 |
| Annealing | 57°C | 90 seconds |  |
| Extension | 72°C | 90 seconds |  |
| Final Extension | 72°C | 10 minutes | 1 |
| Hold | 4°C | ∞ |  |

Pre-Amp Dilution

•	Dilute PreAmp products (1:20) the day of the chip run.

  - Add 3 ul of PreAmp DNA into a dilution plate (Use cheap plate) with 57 ul of Low TE.

  - Pipette 2.5ul/ well if dilution into “Sample Plate”. Do NOT transfer any DNA to NTC wells, just add 2.5 ul of Low TE to “Sample Plate” instead.

Assay Loading Protocol

- Prepare assay loading mix and add 4ul to each well of the “Assay Plate”

1X		110X

2X Assay Loading Reagent					2.5 ul		275 ul

dH20					         			1.5 ul________ 165 ul____________									4 ul		440 ul

- Add 1 ul/well of the assay primer mix from the prepared Assay Stock Plate to have a total of 5ul/well in the final  “Assay Plate”

- Spin down plate. Pipette 4.3 ul of final assay mix into the left side of the Fluidigm Chip

*At this point, prime chip. This will take ~18 minutes

Sample Loading Protocol

- Prepare sample loading mix and add 3.5ul to each well of the “Sample Plate”

1X			110X

Biotium 2X MM							3 ul			330 ul

20X SNPtype Sample Loading Reagent				0.3 ul			33 ul

60X SNPtype Reagent (Purple)					0.1 ul			11 ul

ROX (50X) (Pink)						0.036 ul		4 ul

H2O								0.064 ul		7 ul__________

3.5 ul			385 ul

- Add 2.5 ul of DNA from the diluted sample plate to have a total of 6 ul/ well in the final “Sample Plate”

- Spin down plate. Pipette 5.3 ul of final assay mix into the right side of the Fluidigm Chip

*Chip is now ready to be loaded on IFC controller

Fluidigm Chip Protocol

Priming a Chip

Materials:

- Chip plate

- Syringes

- Unwrap chip plate, save box (note barcode #), check for scratches

- Pull back on syringe and remove stopper carefully to avoid any oil from ejecting

- Insert syringe into holes on either side of the chip so that the inner black bulb is pressed down, begin to inject oil, repeat with new syringe on other side

- Remove blue sticker on underside

- Insert chip into controller aligning angled corner with A1

- Prime ⇒ User ⇒ Run Script

*This takes about 18 mins

After a Chip is Primed

- Eject the chip plate and take to Post-PCR room for loading with Assays and Samples

- Interface Cleaning message: pull red part out and wipe with alcohol wipes

- System Cleaning message: load a blank plastic tray from drawer beneath machines, follow directions on screen

Loading a Chip

Materials:

- Assay Plate

- Sample Plate

- Black Chip Coaster

- Load Assays into left side of chip plate with P10 manual 8 channel pipet set to 4.3 ul

- Stop dispensing at the first stop, you can touch the sides

- When you get to 7A move one row down for the remaining columns to offset sample loading

- Load Samples into the right side of the chip plate with pipet set to 5.3 ul

- Check for bubbles, use universal cheap pipette tips to remove them by hand

- Load chip plate into a Fluidigm controller with angled corner aligned to A1: Load Mix

- ⇒ Run Script

*This will take about 1 hr 30 mins

When Program Finishes, Transfer Chip to Fluidigm ThermoCycler

- Start ⇒ Continue ⇒ use arrow to scroll to SNPtype 96x96 v1

*This will take about 2 hours

Reading the Chip

** If chip absolutely can NOT be read that day (i.e. broken lamp), wrap in wet paper towel and store in fridge. One chip can be read multiple times

- Turn on EP1 Program to warm it up (blue icon) (this takes 5 minutes)

- Use scotch tape to clean Chip carefully

- Insert Chip into large camera machine with angled corner aligned to A1

- EP1 Software protocol:

  - Start New Run

  - Use barcode, unclick to change to “Species_Plate#_Panel#_date”

  - Probe 1 = SNPtype-FAM, Probe 2 = SNPtype-HEX

  - Saved in Runs folder on Desktop

  - All the rest of the defaults

  - Start Run (this takes 3-4 minutes to read)

  - Eject

  - Label Chip plate (ie. TXXX SNPtype) and color in two dots

- Fluidigm SNP Software (green icon)

  - Image View, ROX, 1, Zoom in and change contrast to make sure red outline is around the white glowing spots. If there is a bubble in the spot, select, right click, go to SNP, invalidate SNP.

  - Sample Set-up: New, OK, double click on wells you want to change to NTC;

  - Mapping, M96…

  - Assay Set-up: New, OK, Mapping, M96…

  - Details View: Analyze, Expand too see all SNP results to check

  - Save!

- Log all info in your notebook

- Turn off EP1 software to let lamp cool down

- Turn off computer over the weekend. This will turn off the camera. If computer is off wait at least 30 minutes before imaging a chip

Erasing calls/scores and starting from scratch on Fluidigm chip files:

1. Make a copy of the entire run folder

2. Delete the ChipRun.Processed.bin and ChipRun.bml files

3. Rename the ChipRun.bml.orig file to ChipRun.bml

4. Open the ChipRun.bml file in the appropriate Analysis software

Material and Supplies

| Reagent | Vendor | Catalog Number | Link |
| --- | --- | --- | --- |
| Qiagen 2X PCR MM (100x50ul rxns) | Qiagen | 206143 |  |
| Biotium 2X MM (Fast Probe qPCR) |  | 31005-1 |  |
| 2X Assay Loading Reagent | Fluidigm | BMK-M10-96.96GT-SNP | https://www.fluidigm.com/reagents/genotyping |
| 20X SNPtype Sample Loading Reagent |  |  |  |
| 60X SNPtype Reagent (Purple) |  |  |  |
| ROX (50X) (Pink) | Affymetrix | 75768 | http://www.affymetrix.com/estore/catalog/131187/USB/ROX+Passive+Reference+Dye#1_1 |

| Consumables | Vendor | Catalog Number | Link |
| --- | --- | --- | --- |
| Cheap PCR plates | Fisher |  |  |
| Rainin Multichannel LTS Tips |  |  |  |

Equipment:

- KingFisher 96- for feather extractions

- Fluidigm EP1