How to use the Autoclave

- Place samples on pan spaced apart. If autoclaving liquids, use the non-perforated pan

- Turn the green switch to “START”

- Turn bottom knob to “FILL WATER”

- Watch the water inside the bottom of the autoclave cavity. When the water reaches the front of the door (There is a small notch), turn the knob to “STE”

- Close door and tighten

- Set temperature to desired setting

  - 273F for plastics and other solids

  - 250F for liquids

- Set time

  - 45 mins for plastics and solids

  - 60 mins for Liquids

  - NOTE: SUBTRACT 10 MINS IF THE AUTOCLAVE IS WARMED UP!!!

- The orange “HEAT” light should be on

After the “Ding”

- If you would like the autoclave to dry the samples

  - Wait for the pressure to reach 0 psi

  - Loosen the door and open ~1in to let steam out

  - Set bottom knob to “EXH + DRY”

  - Set timer for ~20-30mins

  - The orange “DRY” light should be on

  - When the cycle is done, set knob to “0” position

  - Open door and remove samples

- If you do NOT want the autoclave to dry the samples

  - Immediately turn the bottom knob to “EXH + DRY”

  - Wait until the pressure reaches 0 psi

  - Turn bottom knob to “0” position

  - Open door and remove samples

- Liquids

  - Wait until the pressure reaches 0 psi

  - Set the bottom knob to the “0” position

  - Open door and remove samples

Once you are done make sure that the bottom knob is set to the “0” position                                                                    and the green switch is set to “STOP”